# Amptek ASCII Configuration File Format

The Amptek ASCII Configuration File Format describes how Amptek applications save and read ASCII command configurations.  ASCII commands are sent in the ‘Text Configuration’ Request Packet (PID1=0x20, PID2=2).  The commands can be sent singly, but most often, they are grouped together into a single "Text Configuration" packet so that the entire configuration can be sent at once.

Amptek Digital Pulse Processor (DPP) and Multichannel Analyzer (MCA) devices using ASCII commands

include DP5(FW6), X-123(FW6), PX5, DPG, MCA8000D, and GAMMA-RAD5.

The number of bytes sent/received in a "Text Configuration" packet must not exceed the maximum size of the corresponding buffer.  If the number of bytes exceeds the buffer size the "Text Configuration" packet must be split into multiple packets. (See the DP5 Programmer's Guide.)

## Configuration File General Rules

- No whitespace before commands

	Commands must start at left column

	Commands must end with semicolon

	Anything after semicolon on a line is ignored (i.e. comments)

	All commands are in upper case

## Configuration File Format

### ASCII Commands and Values

Every ASCII command has a name and a value, delimited by an equals sign (=).   The command name appears to the left of the equals sign. ASCII commands are always the 4-character Mnemonic in upper case that represents the ASCII command.  If no value is present between the equals sign and the semicolon terminator the command is ignored.

Example ASCII command and value:

MCAC=1024;    MCAC - Selects the Number of MCA Channels

### Sections

Commands are grouped into named sections. The section name appears on a line by itself, in square brackets ([ and ]). All keys after the section declaration are associated with that section. There is no explicit "end of section" delimiter. Sections end at the next section declaration, or the end of the file. Sections may not be nested. Section names may be case sensitive depending upon the implementation.

Example Section Names:

[DP5 Configuration File]

[DP5 Configuration Values]

[DP5 SCA Configuration]

### Comments

Semicolons (;) at the beginning of the line indicate a comment line. Comment lines are ignored.  Semicolons (;) at the end of the command indicates where the command ends and optional comments may be placed.  Comments are ignored from after the semicolon up to the end of the line.

Example commented line:

; comment line text

### Blank Lines

Blank lines are ignored with the exception of the "Advanced Programming" example application.

### Repeated Section Names or Commands

All section names are unique and declared only once.  All commands that belong in a section will only be recognized if it is in the corresponding section.  Commands are only allowed once per section, with the exception of the "Advanced Programming" example application.

## Configuration File Sections

There are three configuration file sections:

1.	[DP5 Configuration File] – Stores ASCII commands.

2.	[DP5 Configuration Values] – Stores ASCII command alternate values.

3.	[DP5 SCA Configuration] – Stores SCA register setup values.

### DP5 Configuration File Section

This section is required.  This is the main ASCII command storage section. All commands are stored here except for indexed SCA settings.

### DP5 Configuration Values Section

This section is optional.  This section stores values used to setup configuration dialogs.  Some values are predefined (discrete) while other values are selected from a range (continuous).  Some hybrid commands have both discrete and continuous values.  For hybrid commands it is useful to store the continuous values.  This allows for seamless switching between discrete and continuous values.

### DP5 SCA Configuration Section

This section is required only if SCA counters are being used.  This section is only used by DppMCA and DP5 SDK applications.

### Using/Setting DP5 SCA Configuration Section Values

There are four SCA ASCII commands that do not adhere to "INI File" acceptable rules.  The same    commands are repeated when sent to an Amptek device.  Alternative (INI File compatible) storage must be used to take advantage of widely used (cross platform) "INI File" libraries.  The four commands are:

	SCAH - Set SCA High Threshold

	SCAL - Set SCA Low Threshold

	SCAO - Select SCA Output Level

	SCAI - Set SCA Index

SCAI is used in conjunction with the SCAL, SCAH and SCAO commands to specify which SCA these commands will apply to. When the SCA index is set to 1 through 8, it also enables the specified SCA output and sets it to active high, so the SCAO command doesn’t need to be specified unless a different output option is needed.

The SCA index is stored with the SCA configuration command in the DP5 SCA Configuration section.

| So: SCAI=4;SCAO=OFF;SCAL=1;SCAH=8192; Becomes: SCAO4=OFF; SCAL4=1; SCAH4=8192; | And: SCAO6=OFF; SCAL6=1; SCAH6=8192; Becomes: SCAI=6;SCAO=OFF;SCAL=1;SCAH=8192; |
| --- | --- |

## Using Configurations with the "Advanced Programming" Visual Basic Application

The "Advanced Programming" example application only uses the " DP5 Configuration File" section.  When using the  Advanced Programming" example application, other sections must be removed.

The "Advanced Programming" example application uses repeated SCAI, SCAL, SCAH and SCAO commands in the same order that they are sent out.  (This is not compatible with many "INI File" libraries).

So:

SCAI=4;SCAO=OFF;SCAL=1;SCAH=8192;

Becomes:

SCAI=4;

SCAO=OFF;

SCAL=1;

SCAH=8192;

And:

SCAI=6;

SCAO=OFF;

SCAL=1;

SCAH=8192;

Becomes:

SCAI=6;SCAO=OFF;SCAL=1;SCAH=8192;

# Example DppMCA ASCII Configuration File

[DP5 Configuration File]

RESC=YES;          Reset Configuration

CLCK=20;           20MHz/80MHz

TPEA=12.800;       Peaking Time

GAIF=0.9562;       Fine Gain

GAIN=66.738;       Total Gain (Analog * Fine)

RESL=3276;         Detector Reset Lockout

TFLA=0.200;        Flat Top

TPFA=1600;         Fast Channel Peaking Time

PURE=OFF;          PUR Interval On/Off

RTDE=OFF;          RTD On/Off

MCAS=NORM;         MCA Source

MCAC=1024;         MCA/MCS Channels

SOFF=OFF;          Set Spectrum Offset

AINP=NEG;          Analog Input Pos/Neg

INOF=DEF;          Input Offset

GAIA=20;           Analog Gain Index

CUSP=0;            Non-Trapezoidal Shaping

PDMD=NORM;         Peak Detect Mode (Min/Max)

THSL=3.613;        Slow Threshold

TLLD=OFF;          LLD Threshold

THFA=99.81;        Fast Threshold

DACO=SHAPED;       DAC Output

DACF=50;           DAC Offset

RTDS=603;          RTD Sensitivity

RTDT=8.98;         RTD Threshold

BLRM=1;            BLR Mode

BLRD=3;            BLR Down Correction

BLRU=0;            BLR Up Correction

GATE=OFF;          Gate Control

AUO1=ICR;          AUX_OUT Selection

PRET=OFF;          Preset Time

PRER=OFF;          Preset Real Time

PREC=OFF;          Preset Counts

PRCL=1;            Preset Counts Low Threshold

PRCH=8191;         Preset Counts High Threshold

HVSE=160;          HV Set

TECS=220;          TEC Set

PAPZ=OFF;          Pole-Zero

PAPS=8.5;          Preamp 8.5/5 (N/A)

SCOE=FALLING;      Scope Trigger Edge

SCOT=12;           Scope Trigger Position

SCOG=1;            Digital Scope Gain

MCSL=1;            MCS Low Threshold

MCSH=8191;         MCS High Threshold

MCST=0.00;         MCS Timebase

AUO2=ICR;          AUX_OUT2 Selection

TPMO=OFF;          Test Pulser On/Off

GPED=RISING;       G.P. Counter Edge

GPIN=AUX1;         G.P. Counter Input

GPME=OFF;          G.P. Counter Uses MCA_EN?

GPGA=OFF;          G.P. Counter Uses GATE?

GPMC=OFF;          G.P. Counter Cleared With MCA Counters?

MCAE=ON;           MCA/MCS Enable

VOLU=OFF;          Speaker On/Off

CON1=DAC;          Connector 1

CON2=AUXIN2;       Connector 2

BOOT=ON;           Turn Supplies On/Off At Power Up

SCAW=100;          SCA Pulse Width (Not Indexed - SCA1-8)

[DP5 Configuration Values]

TPEA=12.800;

GAIF=0.9562;

GAIN=66.738;

RESL=3276;

TFLA=0.200;

SOFF=;

INOF=;

CUSP=0;

THSL=3.613;

TLLD=;

THFA=99.81;

DACO=;

DACF=50;

RTDS=603;

RTDT=8.98;

BLRD=3;

BLRU=0;

AUO1=;

PRET=;

PRER=;

PREC=;

PRCL=1;

PRCH=8191;

HVSE=160;

TECS=220;

PAPZ=;

MCSL=1;

MCSH=8191;

MCST=0.00;

AUO2=;

GPIN=;

[DP5 SCA Configuration]

SCAO1=OFF;

SCAL1=0;

SCAH1=1023;

SCAO2=OFF;

SCAL2=0;

SCAH2=1023;

SCAO3=OFF;

SCAL3=0;

SCAH3=1023;

SCAO4=OFF;

SCAL4=0;

SCAH4=1023;

SCAO5=OFF;

SCAL5=0;

SCAH5=1023;

SCAO6=OFF;

SCAL6=0;

SCAH6=1023;

SCAO7=OFF;

SCAL7=0;

SCAH7=1023;

SCAO8=OFF;

SCAL8=0;

SCAH8=1023;