### Amptek Spectrum File Format

DPPMCA saves data in an ASCII format, making it easy to read into other programs. There are several major sections to each file:

- A header, with data including the type of processor and its serial number, the number of channels, start date and time, elapsed live time and real time, text notes, etc.

- Two sections containing information on regions of interest (ROIs) and energy calibration.

- A data section, which is a channel by channel listing of the number of counts.

- A configuration section, listing all the parameters used to configure the processor.

- A status section, showing a readout of critical parameters in the processor, such as the fast and slow counts, the acquisition time, the detector bias voltage, the firmware revision number, etc.

Important notes

- The same file format is used for all Amptek’s processors.  The information contained in the header, the configuration, and status sections depends on signal processor (a digital processor has more parameters than an MCA8000).

- DPPMCA can read files saved in previous releases, e.g the ADMCA software used with the DP4.  If a file saved with DPPMCA is opened in ADMCA, the configuration data will not be recognized.

- The fields are delimited in MS-DOS format (return delimited).  If creating or modifying a file to be read by the DppMCA software using Microsoft Excel, it must be saved as Text (MS-DOS) format.

- The length of the header, ROI, and calibration sections is not a constant but depends on how many ROIs are chosen, the number of points in the energy calibration, how many text notes are entered, and the type of processor.

- The length and values in the configuration and status sections depend on the processor.

File format

| Example | Description |
| --- | --- |
| <<PMCA SPECTRUM>> |  |
| TAG - live_data_1 | TAG is a simple label used to identify the spectrum |
| DESCRIPTION - 55Fe spectrum | Text note |
| <gen> | Header for a “general” note |
| SuperFast SDD | Text note |
| <sys> | Header for a “system” note |
| XR100 with PX5 | Text note |
| <not> |  |
| 142 eV | Text note |
| GAIN - 3 | Indicates number of channels in the spectrum (sometimes called the ADC gain). 0→256 channels, 1→512, 2→1024, 3→2048, etc. |
| THRESHOLD - 0 | Used only for MCA8000A data files. |
| LIVE_MODE - 0 | Used only for MCA8000A data files. |
| PRESET_TIME - 0 | Used only for MCA8000A data files. |
| LIVE_TIME - 53122.498000 | See important note below. |
| REAL_TIME - 53224.315000 | See important note below |
| START_TIME - 06/07/2012 17:14:20 | Time and date stamp when acquisition started. |
| SERIAL_NUMBER - 0 | Used only for MCA8000A data files. |
| <<CALIBRATION>> | Header for calibration section |
| LABEL - eV | Label for units |
| 288.50 1487 | These rows list the data used to produce the calibration. First number is the channel number for a calibration point. Second number is the energy. |
| 1105.26 5895 |  |
| 1214.31 6494 |  |
| <<ROI>> |  |
| 241 273 | These rows list the boundaries of the ROIs. The two numbers are the lower level and upper level of each ROI, in units of channels. |
| 1042 1153 |  |
| 1169 1255 |  |
| <<DATA>> | Header for the data |
| 0 | Counts in channel 0 |
| 0 | Counts in channel 1 |
| 0 | Counts in channel 2 |
| 102156 | Counts in channel 3 |
| 23 | Counts in channel 2045 |
| 21 | Counts in channel 2046 |
| 22 | Counts in channel 2047 |
| <<END>> | Field indicating the end of the data section |
| <<DP5 CONFIGURATION>> | Header for configuration section |
| RESC=?; Reset Configuration | These rows list all of the parameters which have been sent to the digital processor, to command its status. For their meanings, refer to the Programmer’s Guide. The parameters and allowed values vary with the processor. |
| CLCK=80; 20MHz/80MHz |  |
| TPEA=1.000; Peaking Time |  |
| GAIF=1.0573; |  |
| GAIN=29.998; Total Gain (Analog * Fine) |  |
| <<DP5 CONFIGURATION END>> | Field indicating the end of the configuration section |
| <<DPP STATUS>> | Header for status section |
| Device Type: PX5 | These rows list parameters read back from the processor, including the fast counts, slow counts, acquisition time, etc. |
| Serial Number: 2079 |  |
| Firmware: 6.06 |  |
| FPGA: 5.14 |  |
| Fast Count: 142748187 |  |
| <<DPP STATUS END>> | Field indicating the end of the status section |

- Real Time:  For the MCA8000A and for the DP5 family of products (DP5, PX5, Gamma-Rad5, X-123s with DP5, etc), this is the real time.  For the DP5 family of products (DP4, PX4, Gamma-Rad), this number is actually the acquisition time

- Live Time:  This number represents the total amount of time the processor is acquiring data.

  - In the case of the MCA8000A, this is the conventional live time of analog electronics, the real time clock is stopped when pulses pile up.  For the MCA8000A, the user should use the ratio of live and dead time to correct the total counts and obtain the corrected input count rate.

  - In the case of digital processors, however, this live time is different.  The real time is stopped when pulses pile-up and also when the system stops taking data, i.e. during the lockout interval following a reset.  The user can use the ratio of live and acquisition time, not real time, to obtain total count rate.  Even better would be to use the directly measured fast counts.  The fast number of counts and the acquisition time are both in the status section of the file.