| 第14卷第2期 |  | Analytical testing techniques and instruments | Vo lume 14 N umber 2 |
| --- | --- | --- | --- |
| 2008年06月 |  | ANALYSIS AND T EST ING TECH NOLOGY AND INST RUMENTS | June 2008 |

大型仪器维修与维护( 125~ 126)

JSM 5610LV 扫描电镜的日常维护

And common troubleshooting

Li Jianping, Yang Yongdong

  - Yantai University, Yantai 264005)

Abstract: In the daily maintenance management of scanning electron microscopes, the environmental factors of the room should be controlled properly, the interior of the mirror should be kept in a high vacuum condition for a long time, and the accessory equipment should be checked regularly to see if it meets the requirements.  Explain how to troubleshoot problems such as when the mirror always freezes, and the black screen shows an abnormal image.

Keywords: Scanning electron microscopes;  Routine maintenance management;   Troubleshooting

| 中图分类号: T N 16 | 文献标识码: A | 文章编号: 1006 3757( 2008) 02 0125 02 |  |
| --- | --- | --- | --- |
| With the development of science and technology, the use of electron microscopy has increased |  | (6) regularly check the condition of the water circulation and the water level below the line of fire |  |
| The more widespread it is, it has contributed to the technological revolution and, in particular, to the development of nanotechnology |  | Add purified water immediately. If the circulating water is cloudy and dehydrated, replace it immediately |  |
| The electronic microscope is sophisticated and expensive |  | To prevent clogging and fouling of the circulatory system, which may affect the water circulation system |  |
| This requires specialized management, use, maintenance knowledge and assurance |  | Productivity. |  |
| In normal working conditions, better serve scientific research and production. This article |  | (7) periodically switch on the backscatter electron imaging system, which is not commonly used for electromirrors |  |
| 针对日本电子的 JSM 5610LV 扫描电镜, 简谈一下 |  | Systems, Component Imaging Systems, Topographic Imaging Systems, Stereographic Imaging Systems, etc. |  |
| Daily maintenance and troubleshooting of common problems. |  | Function system to prevent aging of electrical components. There should be no less time at each turn on |  |
| 1 日常维护 |  | 于 1 h. |  |
|  |  | ( 8) 定期检测电镜的 X 射线是否泄漏, 加强防 |  |

- 1)keep the mirror room clean and dust free, use hand protection such as air conditioning and humidity, and ensure personal safety.

| 段, 控制好室内环境因素, 使室内温度在15~ 25 | 2 常见故障的排除 |  |
| --- | --- | --- |
| 之间, 湿度在50% 以下. |  |  |
| (2) keep the mirror in a certain vacuum for a long time | When a mirror fails, the main part of the fault should be determined first |  |
| 一般要求在 10- 5 T orr 以上[ 1] . Even the electroglass is not working | Bit or cause, then step by step to maximize the extent of the fault |  |
| 期间, 也要定期抽真空, 雨季里 5 d 抽一次, 其它季 | The problem was eventually identified as the cause of the failure or the specific component, and the question was then asked |  |
| This should be done once to prevent the electromirror cartridge and the internal components of the spectrometer | 题解决. 以下是日本电子的 JSM 5610LV 扫描电镜 |  |
| Rust. | There are several common troubleshooting methods. |  |
| ( 3) 经常检查机械泵( DP) 的油液面, 看其液面 | 2.1 故障1 |  |
| Is it at the level of the window level engraved, if it is at the window level engraved | Problem: The screen shows a picture, but the picture is not jittery and blurry |  |
| If mechanical pump is observed, add mechanical pump oil immediately | Clear, and the upper and lower interrows are not clear. |  |
| 窗口油呈茶色, 应马上更换机械泵( DP) 油. | Cause: There are many reasons why the image is not clear, in this case |  |

- 4)定期更换电镜内的扩散泵(RP)油[2] .Dynamic focus balancing was not enabled during operation

| (5) regularly clean the moving diaphragm, sample stage, and electrons inside the barrel | ( Dynamic Focus Cor rection) 功能, 电镜产生交变 |
| --- | --- |
| The gun room is sensitive to contamination of parts, ensuring that the electronic light path works properly. In the | 磁场, 使焦距抖动, 致图像抖动、模糊不清[ 3] , 且 |
| During cleaning, the diaphragm is heated with a vacuum sprayer, the sample stage and the | The upper and lower interrows are not clear. |
| 电子枪室用棉球蘸酒精、丙酮1 1 混合液擦洗. | Solution: Turn off dynamic focus balancing. |

收稿日期: 2008 03 14;   修订日期: 2008 04 16.

作者简介: 李剑平( 1963- ) , 男, 大专, 实验师,  主要研究方向扫描电子显微镜.

126	Analytical testing techniques and instruments 	第 14 卷

2.2	故障2

Problem: The phone freezes, the focus and magnification keys all fail, and they don't work, especially the tungsten lamp in the mirror, the power switch of the Oh/Ready cannot be switched off.

Cause: The adjustment amplification rate is too fast and the operation is incorrect, the call mirror related procedure is not smooth.

Solution: Switch off the mirror and restart the mirror.

2. 3故障3

Problem: When the mirror is switched on, the screen does not display a picture! ! ! Blank screen.

Cause: The filament is broken and does not power on or the brightness and contrast of the screen display are not appropriate [ 4 ].

解决方法: 首先打开电子枪( Gun) 图标, 检查钨灯丝好坏[ 5] ( 是否通电) , 若不通电, 说明钨灯丝烧

If so, replace the filament;  If power is applied, adjust the contrast and brightness until the image is displayed and clear.

2.4  故障4

故障现象: 电镜样品台Z 轴显示值为零, AFC、ACD、A ST 等自动功能无法实施.

原因:  电镜的 EOS 电路板的 A/ D 转换器

M P8775AN 功能异常, 无法实施自动功能, Z senso r

No output.

解决方法: 更换电镜的EOS 电路板的 A/ D 转

换器MP8775AN.

2.5  故障5

故障现象: 电镜在实施扫描过程中, 当 SPOT SIZE 数值设置是16 时, 显示的图象正常; 当SPOT SIZE 数值设置大于17 时, 显示的图象抖动, 并且伴

Discharge phenomenon.

Cause: The tungsten filament of the mirror is rapidly burned and appears to be in a state of constant flux, the current loss adjustment of the tungsten filament is adjusted by [ 6 ], thus triggering the unstable electron. When the spot size value is set to greater than 17, the primary magnetic lens is unstable in converging the electron beam, which affects the imaging effect.

Solution: Change the tungsten filament of the mirror, and make a series of adjustments to the medium range.

- Conclusion

Having a high degree of responsibility, doing a good job in the daily maintenance of the mirror, having certain troubleshooting skills, is a basic requirement for the mirror operators.  Only by doing this can we better perform the role of electroscopy for scientific research.

References:

- 1 ]   马金鑫 , 朱国凯. 扫描电子显微镜入门[ M ] . 北京:  科

学出版社, 1985. 62.

- 2 ]   朱宇. 电子显微镜的日常维护及故障的检修[ J] .  内

蒙古石油化工 , 2000, ( 3) : 83.

- 3 ]   朱祖福 . 电子显微镜[M ]. 北京: 机械工业出版社,

1984.  56.

- 4 ] 李剑平. JSM 5610LV 扫描电镜的一般故障排除[ J] . 仪器仪表用户, 2007, 14( 3) : 146.

- 5 ]   王化正 , 李玉生 . 现代分析仪器维修手册[ M ] . 北京:

航空工业出版社, 1988. 415.

- 6 ] 刘跃华 , 徐彦平. JSM T 300 扫描电镜维修四例[ J] . 电子显微学报, 1998, 17( 5) : 672 673.

Daily Maintenance and Familiar Faults Clearance of JSM 5610LV

LI Jian ping , YANG Yong dong

( Institute of Science and T echno logy

for Opto E lectr on Inform ation, Yantai U niversity , Yantai 264005, China)

Abstract: During the daily maintenance and management of SEM , w e should better control the factors o f the indoor environmental and keep the interior of SEM at a high vacuum state and check the ancillary equipment w hether they reach the demands or not termly. Finally, the metho ds of fault clearance such as halt, black scr een and imag e show abnormality of SEM is introduced in this paper.

Key words: SEM ; maintenance and management; fault clearance

Classifying number: O657. 32