OOISesnors with MFPF
OOISensors was updated to include support for TauTheta’s MFPF in addition to all of Ocean Optics’ spectrometers.  Default, first time start-up configuration depends on software version.  Switching is possible in the software configuration panel (Configure | Software) top-right corner control named “System Type”.  Application front panel and configuration window will change to display different components according to hardware configuration.
OOISensors for spectrometers requires the application to be configured for “Spectrometer” as the System Type in the software configuration panel.  It also requires application to be configured for the connected spectrometer using the hardware configuration panel (Configure | Hardware).  In some situations user might need to configure this manually.  To configure manually, make sure application is not running and edit the OOIDrv.INI file located in the Windows folder to specify the correct SpectrometerType, ADCType and “Initialized=1”.  This change will not take effect if OOISenosrs or any of the other software from Ocean Optics is running at the time of change.
OOISensors for MFPF requires the application to be configured for “MFPF” as the System Type in the software configuration panel.  It also requires two additional software packages which set-up the USB port as a serial COM port and communicate with the TauTheta device.  The following two software packages are needed before MFPF would work:
1. USB installation which actually emulates the USB interface as a COM port.  After installation, check the device manager to see which COM port was actually assigned to it.  Run PreInstaller.exe 
2. MFPF host DDE server.  Run Setup.exe
3. Connect the TauTeta device only after installing these two items. 
The first, USB installation, is requires to be installed only once at the beginning.  Once this software is installed and a device is connected, check the device manager for the COM port that was assigned to the MFPF device which is actually connected to a USB port.  Once the USB software is installed, the host DDE program can be installed.  Normal sequence of operation should be to run the host program and then start OOISensors application.  If OOISensors application is started without the host program already running, an error message will tell you that it could not find the DDE host.  If that happens, quit OOISensors and start the host program before starting OOISensors again.
Change MFPF parameters using the host program.  OOISensors cannot control the device.  OOISensors does monitor for channel configuration changes and will reconfigure display accordingly.
Switching between spectrometers and MFPF will remember most configuration parameters as hardware specific but some parameters are common to all hardware types.  When switching from MFPF to spectrometers, the spectral graph will not be visible.  To see the spectral graph, go to the software configuration panel and select the “Display” tab.  Then change the “Graph and Chart Display Mode”.  

