Solartron Instruments 1887/1988

Specification issue D ?

7081 noise :

0.5 ppm p-p (5 V)  0.09 ppm rms

Transfer noise is 0.3 ppm p-p, 24 hrs instability is 1.2 ppm.

Temp-co is 0.8 ppm/C at +/–5 C.

Solartron 7071/7081 Operating manual, 1989

Specification issue E ?

Transfer noise is the same 0.3 ppm p-p,

but 24 hrs instability is 0.5 ppm now!

Temp-co is lower (0.5 ppm/C), but temperature span is lower too (+/–3 C).

Instruments test specification, 1994

Allowable settling error is 2.6 ppm. I. e. readings discontinuities can be as large as 26 V!

Allowable 10 V noise is 3 V rms that approximately equal to 18 V p-p!

Solartron 7081 statistics by Lymex Zhang (lymex / bg2vo)

DMM’s low frequency noise, Standard deviation

Measured 7081 rms noise at 7 ½ mode is 0.095 ppm (1 V rms or 6 V p-p).

Measured 7081 rms noise at 8 ½ mode is 0.062 ppm (0.6 V rms or 4 V p-p)