ViVA Help Documentation

Table of Contents

I. INTRODUCTION	4

I-1. Version	4

I-2. Usage Guidelines	4

I-3. System Requirements	4

I-4. Installation	5

I-5 Image Acquisition Versus Image Display	5

I-6 User Interface Design	5

I-6-1 Image Window Types	6

I-6-2 Image Selection Indicator	7

I-6-3 Image Layouts and Other View Options	7

I-6-4. Toolbars	9

I-7. Common Semantics	9

II. Open/Close/Save/Print Copy Images	10

II-1. Supported File Formats	10

II-2. Open Image	10

II-3. Close Image	10

II-4. Save Image	11

II-5. Print Image	11

II-6. Copy Image	11

III. Image Manipulation	12

III-1. Cursor Functions	12

III-1-1. Standard Selection Cursor 	12

III-1-2. Zoom Cursor 	12

III-1-3. Magic Window Cursor 	13

III-1-4. Pan Cursor 	13

III-1-5. Region of Interest Cursor 	13

III-1-6. Mouse and Key Shortcuts	14

III-2. Window/Level Functions	15

III-2-1. Window/Level Scroll Bars	15

III-2-2. Edit W/L Button	15

III-2-3. Auto W/L Button	17

III-2-4. Invert W/L Checkbox	17

IV. Image Acquisition	18

IV-1. Open/Close/Check A Communication Link	18

IV-2. Acquisition, Calibration and Other Operations	19

IV-3. System Settings	20

IV-4. Mode Settings	22

V. Preferences	26

VI. Physics/Engineering Use	28

VI-1. Display Test Pattern	28

VI-2. Set Debug Verbosity On/Off	28

VI-3  Set Dcds Off	28

VI-4. Transmit Files	28

VI-5. Retrieve Files	29

VI-6. Analog Offset Calibration	30

VI-7. Image Arithmetic	31

VI-8. Find Noisy Pixels	32

VI-9. Defect Map Editor	34

VI-9-1. Editing A Pixel	35

VI-9-2. Editing A Row Segment	36

VI-9-3. Editing a Column Segment	36

VI-9-4. OR Image Operation	37

VI-9-5. Retrieve Images	37

VI-9-6. Transmit Image	37

VI-9-7. Close Defect Map Editor	37

VII. Statistics	38

Navigation within this document is assisted with bookmark jumps, which are indicated as underlined text.  Clicking on a section title in the index listing will move the cursor to the beginning of the section.  Vice versa, click the section title at the beginning of the section to move the cursor back the section title in the index listing.

# I. INTRODUCTION

## I-1. Version

This version of ViVA (Varian Image Viewing and Acquisition) is 2.0; the revision letter is J; and the build number is 28. This information is available at the Help|About ViVA command, which displays the dialog box in the following figure.

Figure 1: The About ViVA dialog box.

If a link is open to the command processor at the time the dialog box is opened, the Command Processor is queried for Version information as in the figure.

## I-2. Usage Guidelines

ViVA is provided and intended for research and development purposes only.

## I-3. System Requirements

ViVA will run on Intel-based PCs using Windows 95, Windows 98, or Windows NT 4.0.  It has been tested with the following video settings: 8-bit (256 colors), 16-bit (65536 colors), 24-bit (16 M colors), and 32-bit (sometimes called "True Color").

Acceptable performance in image viewing and manipulation can be obtained with a 166 MHz Pentium CPU system, or faster. The number of open images is limited by the combination of physical and virtual memory. Physical memory of 64 MB or more is recommended.  The quality of the video card hardware, firmware, and driver software also impacts the perceived speed of operation.

## I-4. Installation

The Setup.exe in the root directory of the PaxScan CD provides an automated installation process for ViVA and its dependent files.  It also creates a shortcut of ViVA in the desktop.  Setup.exe is automatically started when the CD is inserted onto a CD drive unless the Auto-Run CD option is turned off.  To start Setup.exe manually, use the Run command under the Start button to execute X:\Setup.exe, where X is the drive letter of the CD drive that is containing the PaxScan CD.

To uninstall ViVA, run Add/Remove Program application in the Control Panel and remove the entry PaxScan.  Note that running this automatic uninstall program will also remove all other PaxScan components such as Command Processor files and Hostdown files that have previously installed onto the hard drive.

A description of the features of ViVA can be found in this help file ViVA_Help.doc.  Viewing this help file requires Microsoft Word application.  If you need to obtain a copy of Microsoft Word for this purpose, you can download a copy of MS Word Viewer 97 on http://www.microsoft.com/Office/office/000/viewers.htm.  Then install it onto your computer before viewing ViVA_Help.doc.

To run ViVA, locate and double click the shortcut of ViVA on the desktop.

## I-5 Image Acquisition Versus Image Display

ViVA can be used to control the VIP-9 flat panel X-ray imaging system from a Windows PC.  ViVA can also be used to display and manipulate images on a Windows PC that is not connected to a VIP-9 system.

## I-6 User Interface Design

A screen shot of the ViVA user interface is shown in the following figure.

Figure 2: ViVA screen shot

### I-6-1 Image Window Types

ViVA does not use the standard Windows MDI (Multiple Document Interface) arrangement of multiple, overlapping, moveable and resizable windows.  Instead, it offers two other types of windows for displaying images: a scrolling image-gallery or Thumbnail View window and fixed frame Image View windows.

Newly opened or acquired images appear in miniature in a Thumbnail View window.  This scrolling view holds all the active and opened images.  The Thumbnail View window can be positioned horizontally at the top of the application's main window or vertically along the left edge.

Images can be moved by drag and drop to one of the larger Image View windows.  There may be one, two or four of these windows.  Their position is fixed within the application's main window frame.  Most image manipulation operations can only be performed on an Image View and not on a Thumbnail View.  In addition to the image file name, this window also displays information about magnification (zoom), window/leveling, image size, and, for the pixel under the mouse cursor, (x, y, value) data.

The size of the Image View windows is adjusted automatically to fit the size of the application’s main window.

The arrangement and number of windows can be selected using the View pull-down menu.

### I-6-2 Image Selection Indicator

When an image is dragged from the Thumbnail View to one or more Image View windows, the separate views of the image are logically linked.  Only one copy of the image is stored in memory for use by all views of that image.

A Thumbnail View or Image View can be selected by clicking on the view with the mouse pointer.  The currently selected image view is indicated with a bright green selection rectangle.

Other linked views of the same image shown in the selected view are indicated with a bright blue (cyan) selection rectangle.

### I-6-3 Image Layouts and Other View Options

Under the View menu, there are there groups of options.  Their uses are explained as follows:

- The two options Image Thumb on Top and Image Thumb on Left allow the user to place the Thumbnail View of all open images either on the top or on the left of the main window.

- The three options Single-Image Layout, Two-Image Layout and Four-Image Layout allow the user to choose the number of images that can be viewed in the Image View at the same time.  The options are one, two or four images, respectively.

- Image Border allows the user to display or to hide a green border of an image in the Image View. Image Information is applicable to images saved in the standard ViVA format .viv which is new in Rev J. Images retrieved with Rev J software default automatically to saving a file header which contains information on settings in effect when the image was obtained. The Image Information command displays the information contained in this header in a series of tab enabled dialog boxes which are illustrated in the following figures:

Figure 3a: Image Information:Image tab.

Figure 3b: Image Information:System tab

Note that boxes with a white background are for user input. Also the Video tab, although accessible, is present only for future implementation purposes.

- The option, Status Bar, allows the user to display or hide the status bar on the bottom of the main window.  The status bar displays a short description of a pull-down menu command when the command is highlighted by the cursor.  Alternatively, the status bar displays the currently selected mode (Fluoroscopy, Radiography…) when the cursor is placed in other sections of the application.  The other view option, Reset User Messages, allows the user to choose to display all user messages.   These messages are warnings or explanations that are routinely displayed before completion of a command and at which time the user may select not to display the message again.

Figure 4: A user message

### I-6-4. Toolbars

Two toolbars are used in ViVA.  The Image Edit toolbar is oriented vertically and may be docked along either the left or right edge of the main ViVA window.  The Acquisition or Defect Map Editor toolbar is oriented horizontally and can be docked along either the top or bottom edge of the main ViVA window.  The toolbars are capable of "floating" over the main ViVA window.

The functions supplied by these toolbars are described below in Sections III and IV.

## I-7. Common Semantics

ViVA uses a number of common terms and notations that have special meanings.  This section describes the meanings of these terms.

- Menu items ending with three periods (e.g. File|Open…) indicate that executing the commands will require further inputs from the user in a dialog box.  On the contrary, menu items without the periods require no further involvement by the user and will be executed immediately.

- In a dialog box, a Cancel button implies closure of the dialog box without executing data on the dialog box; an Apply button implies execution of data on the dialog box that updates data on disk and/or on the command processor.  An OK button performs the same tasks by the Apply button if they have not been executed, then closes the dialog box.

# II. Open/Close/Save/Print Copy Images

## II-1. Supported File Formats

ViVA now defaults to “viv” format which saves a 2048 byte header containing both optional user entered information as well as automatically retrieved settings in effect at the time an image was obtained. Additional information is given in Section I-6-3. The previous “raw” format (no header, 16 bit unsigned) is still supported, as well as Visilog format (Visilog is a commercially available image processing software application). ViVA can also save images in a Windows bitmap format (BMP 8-bit) and JPEG (Joint Photographic Experts Group) format (JPEG grayscale). Also new in Rev J, if “any” file format is selected the user is prompted to enter information as to the header and image size if known allowing generic file types to be opened for viewing provided they follow a layout of header and followed by pixel values expressed as 16-bit unsigned integers. This has limited scope and cannot be applied to compressed files; header information is ignored and cannot be accessed by the Image Information command. A dialog box illustrated in the following figure will be displayed when opening a file with the “any” type specified.

Figure 5: “Any” file specification dialog.

## II-2. Open Image

Open an image for display and manipulation using the menu option File|Open...  The newly opened image will appear in the Thumbnail View.  ViVA can open images in viv, Visilog, raw formats as well as certain generic file types as described above.

The opened image is first displayed with an automatic window-leveling setting, which is intended to show image features across a broad range of contrast.

Upon a success of opening an image, the path and file name of the image is saved onto the most recently used file list as the newest entry, which can be used as shortcut to open the file.

## II-3. Close Image

This close function consists of two options.

Close an image.  If a Thumbnail View is selected, choosing File|Close will close not only that view of the image, but all associated Image View windows as well (see Section I-6-1 for description of window types and Section I-6-2 for image selections).  If an Image View window is selected, only that view of the image will be closed, and the Thumbnail View of the image, and any other associated Image View windows, will remain open.

Close all currently open images.  Choosing File|Close All will close all open images in both Thumbnail View and Image View windows.

## II-4. Save Image

Save an image.  Choosing File|Save As… will display a standard Windows File Save dialog box.  The user can choose to save an image in viv, raw, Windows Bitmap, JPEG or Visilog format.

## II-5. Print Image

This print function consists of three options.

Print an image.  Choosing File|Print… will print the currently selected image either in the Thumb View or in the Image View.

Display a print preview of an image.  Choosing File|Print Preview… will display a print copy of the currently selected image either in the Thumb View or in the Image View on the screen.  This is intended to show the layout of a printout of an image.  It does not reflect the quality of the printout.

Display printer setup.  Choosing File|Print Setup… will allow the user to select a printer and paper options.

## II-6. Copy Image

Copy the current image into the clipboard.  Choosing Edit|Copy will copy the currently selected image into the clipboard memory.  Note that this function always copies the whole image using the current window-level settings, not just the section of the image being displayed in the image view.

# III. Image Manipulation

The Image Edit (vertical) toolbar supplies controls for the following functions:

- Pixel value display

- View selection

- Zoom in/out

- Pan

- 2x zoom "magic window"

- Window/level (auto, invert, linear/s-curve, scroll bar control...)

## III-1. Cursor Functions

This section describes the five-cursor types currently implemented in ViVA Image Edit toolbar.

For all the cursor types, the (x, y, value) triplet is displayed in an Image View window, which is the image pixel corresponding to the point under the active spot on the cursor.

### III-1-1. Standard Selection Cursor

The standard selection cursor also called point cursor, looks like a crosshairs with a small arrow to the lower right.  It is selected by pressing the button at the upper left of the toolbar.  The active pixel for this cursor is at the center of the crosshairs.

This cursor is used for view selection (click on the view) and for drag and drop of images from view to view.

The operations enabled by the other cursors are also available when the standard selection cursor is active, if the user employs certain key combinations.  These key combinations are described at the end of each section describing the other cursors, and are summarized in Table 1 at the end of Section III-1.

### III-1-2. Zoom Cursor

The zoom cursor looks like a small magnifying glass with plus and minus signs on either side of its handle, superimposed on a rectangle with a vertical and horizontal bar through it.  This image is meant to signify Zoom Window.

When this cursor is selected, the user can zoom in on an image in an Image View window by clicking the left mouse button on the view.  If the newly magnified image is larger than the view window in either direction, the point on which the user clicks will jump to the center of the window.

The user can zoom out by clicking the right mouse button.

ViVA only zooms in factors of two.  This is meant to be a feature, which maintains the most direct relationship possible between the image pixels and the screen pixels.  The range of magnification is from 1/128x to 8x.

This same functionality is available when the standard selection cursor is in use and the user holds down the Shift key while clicking.

### III-1-3. Magic Window Cursor

The magic window cursor looks like a small magnifying glass with a plus sign on the left side of its handle.  This image is meant to signify Magic Window.

When this cursor is selected, clicking and holding down the left mouse button on an Image View window will display a moveable window in which the same image is viewed at twice the underlying resolution.

This same functionality is available when the standard selection cursor is in use and the user holds down the Control key while clicking with the right mouse button.

### III-1-4. Pan Cursor

The pan cursor looks like two crossed and double-headed arrows, one vertical and one horizontal.

When this cursor is selected, clicking and dragging with the left mouse button on an Image View window will cause the image to move around (i.e., pan) within the window.  If the image is displayed with a magnification that results in the image being smaller than the window in one direction or the other, then the image will not scroll in that direction.

This same functionality is available when the standard selection cursor is in use and the user holds down the Control key while clicking with the left mouse button.

### III-1-5. Region of Interest Cursor

The Region of Interest or ROI cursor looks like a crosshairs, which is enclosed in a dashed square, with a small arrow to the lower right.

When this cursor is selected, the ROI mode is activated.  The cursor can then be used to specify a region on an image for statistical calculations.  The options in the Statistics pull-down menu and the ROI Stats button on the Image Edit toolbar also become active.  If a region had been previously specified, it is redisplayed on the image as well.

### III-1-6. Mouse and Key Shortcuts

Shortcut keys or hot keys provided in ViVA are summarized n the following table.

Table 1: Shortcut key combinations for cursor operations, with standard selection cursor active

| Operation | Key Combination |
| --- | --- |
| Selection/drag and drop | Left mouse click |
| Zoom in/out | Shift + left/right click |
| Magic window | Ctrl + right click |
| Pan | Ctrl + left click |
| Specify a region of interest and show statistical information | Ctrl + drag with left mouse button |

In addition, ViVA also provides three context menus that apply on three different types of images: Thumb View images, Image View images in viewing mode and Image View images in Defect Map Editor mode.  To activate a context menu of an image, point the mouse cursor at the image and hold down the right mouse button for at least one second.  The contents of the menus are listed in the following figures:

Figure 6: Context menus for Thumb View images, Image View images in viewing mode and Image View images in Defect Map Editor mode respectively

## III-2. Window/Level Functions

Window and level can be changed in several ways using controls on the Image Edit toolbar.  In this section, the words window and level are used in the sense of adjusting the brightness and contrast of the image in a way that is common for medical imaging applications.

Window/level control is the only form of image manipulation that can be applied to Thumbnail View windows.  Window/level settings are maintained during drag and drop operations.

### III-2-1. Window/Level Scroll Bars

There are two scroll bars in the Image Edit toolbar that can be used to change window and level.  The left scroll bar moves the bottom of the window up and down, while the right scroll moves the top of the window.  Holding down the Shift key while manipulating either scroll bar causes the scroll bar buttons to move up and down together, which corresponds to changing the level.

### III-2-2. Edit W/L Button

Clicking this button opens a dialog box which allows the user to adjust the window/level of the image by manually typing in new settings, either in terms of bottom/top (as with the scroll bars), or in terms of window and level. The dialog box is shown in the following figure.

Figure 7: Edit Window/Level dialog box

The relationship between window/level and bottom/top is:

window = top – bottom,

level = (top + bottom) / 2.

This dialog box also lets the user choose between linear window/level mapping and three non-linear mappings.  Non-linear (Atan/S-curve) was present in earlier versions of Viva while Non-linear(Tanh) and TriLinear are new as of version 2 rev.H.  These functions are applied in cooperation with the setting of bottom and top. Actual pixel values (0-4095) are mapped to a gray level (typically 0-255) stored in a lookup table (LUT).  For the linear setting this mapping effectively results in maximum contrast in the bottom/top region but none outside that range.  The Non-linear (Atan/S-curve) is based upon the arctangent function and an example of the mapping for particular bottom/top values is shown in the following figure together with the linear mapping.

Figure 8: Gray level mappings for linear and Atan/S-curve functions

The Atan function effectively softens the response allowing some contrast outside the bottom/top range, and generally results in a more pleasing image.  The Tanh function behaves very similarly to the Atan but is intermediate in action between the Atan and the linear.  The TriLinear function creates three variable linear responses in the ranges 0--bottom, bottom--top and top--4095, referred to as Lower, Middle and Upper respectively.  The user may customize the response by setting the ratios of the slopes in the dialog box.

White level blanking may be useful for example when only a relatively small part of the image is of interest and some substantial part of it was non-absorbing to X-rays normally resulting in an expansive white area.  When activated, the white area blanking allows the user to specify a “threshold” value above which lookup table values are replaced with the replacement value.  The result, in appropriate cases, is an overall image in which the surrounding white areas do not detract from the region of interest.  Note that the net effect of the threshold and replacement values will depend on which window/level mapping is selected as well as the bottom/top values.

### III-2-3. Auto W/L Button

Clicking this button applies that same automatic window/level algorithm that is used when images are opened or acquired.

### III-2-4. Invert W/L Checkbox

Normally, low pixels values are mapped to dark gray levels, and high pixel values to bright (whiter) gray levels.  Checking this box inverts the normal relationship.

# IV. Image Acquisition

The functions in this category are designed to control a command processor and to obtain data from it.  A subset of these functions is placed on the Acquisition toolbar for quick access.

The Acquisition toolbar is shown in the following figure in its floating state.

Figure 9: Acquisition toolbar in floating state, with an Ethernet connection to a command processor

The choices available on the Acquisition menu are:

- System Settings

- Mode Settings

- Offset Calibration

- Gain Calibration

- Acquire Image

- Reset State

- Retrieve Image

- Transmit Image

- Open/Close Ethernet  Link

- Open/Close Serial Link

- Check Ethernet/Serial Link

- Reset Link

## IV-1. Open/Close/Check A Communication Link

ViVA supports an Ethernet or a serial link to a command processor.  Only one link can be established between ViVA and a command processor at a time.  The Open/Close/Check options in the Acquisition menu are dynamically labeled so that only appropriate functions are displayed and functional.  For example, after an Ethernet link is open, the Acquisition|Open Ethernet Link option is changed to Acquisition|Close Ethernet Link; the Acquisition|Open Serial Link option is grayed out and Acquisition|Check Ethernet Link option becomes selectable.

When there is no open connection to a command processor, the Acquisition Mode selection list on the Acquisition toolbar will display << No Connection >> in the box where Fluoroscopy – 30 FPS is displayed as shown in the above figure.  To open a connection, choose either Open Ethernet Link or Open Serial Link option from the Acquisition menu.  Communications settings are available via the Preferences dialog box, described in Section V.

Once a connection is opened, the Acquisition Mode selection list will offer a list of the operational modes available for the command processor.  To select a different operational mode, simply choose the name of the mode from the list.

## IV-2. Acquisition, Calibration and Other Operations

- To initiate image acquisition, click on the Acquire Image button on the toolbar, or choose Acquisition|Acquire Image.  Once the acquisition is complete, the user may capture an image from the command processor by clicking the Retrieve Image button or by choosing Acquisition|Retrieve Image. A window is opened which allows the user to input comments and other technique information. The dialog box is illustrated in the figure below.

Figure 10: Image Info: User input

This information may also be entered or edited later using the View|Information Information command.

- Previously retrieved images may be sent to the command processor for live display.  First, select the appropriate image by clicking on it in the ViVA interface; then click on the Transmit Image button or choose Acquisition|Transmit Image.

- To initiate offset or gain calibration, choose either Acquisition|Offset Calibration or Acquisition|Gain Calibration accordingly.

- The command processor can be reset to a known state by choosing the Acquisition|Reset State option.  This will cause the system to return to the default-operating mode, which is the first mode in the Acquisition Mode selection list.

- To check the status of the communication link between ViVA and the command processor, choose Acquisition|Check Ethernet/Serial Link.  The appropriate word Ethernet or Serial is displayed in the command option as ViVA keeps track of the current link type.

- To terminate the communications link, choose Acquisition|Close Ethernet/Serial Link. The appropriate word Ethernet or Serial is displayed in the command option as ViVA keeps track of the current link type.

- If the connection is lost and the system is not in synchronization with the Command Processor, choose Acquisition|Reset Link to restart the communications without terminating ViVA.  This command will close all links and clear all information obtained from the Command Processor.  Then choose Acquisition|Open Ethernet/Serial Link to establish a link.

IV-3. System Settings

When there is an open communications link to a Command Processor, the user may select the Acquisition|System Settings… option.  One of the two dialog boxes shown in the following two figures will appear.

Figure 11a: Dialog box of System Settings with ABS

Figure 11b: Dialog box of System Settings without ABS option

On the left side of the dialog box, unchangeable information about the system is displayed.  On the right side of the dialog are Image Corrections settings.  These four check boxes allow the user to select which image correction algorithms are applied in the Command Processor.  Please, see the Instruction and Operations Manual of the VIP-9 Fluoroscopic Imaging Products for details about these algorithms.  Note that the Pixel Defect Map option can only be selected if the Gain Corrections option has been selected.  Similarly, the Digital Line Noise Reduction option can only be selected if the Offset Corrections option has been selected.

If the Command Processor has an ABS video board, the dialog box will include the ABS Gain Control option as displayed in Figure 11a, which an additional correction algorithm as well as vertical and horizontal reversal check boxes.

## IV-4. Mode Settings

With an open communications link, the user may select the Acquisition|Mode Settings option.  This will cause the Mode Settings dialog box to appear.  This dialog box is shown in Figures 8a, 8b and 8c, below.

At the top left of the dialog box is a selection list for the operational mode.  The rest of the settings shown in the dialog box apply only to the selected mode.

The Information group of settings shows unchangeable settings, for information purposes.

The Acquisition/Display group shows different settings depending on whether a mode with fluoroscopy (continuous) acquisition type or radiography (accumulation) acquisition type is selected.  For a fluoroscopy type mode, the user can set the recursive filter fraction as shown in Figure 12a.  For a radiography type mode, the user can choose between letting the system automatically sense X-rays or having it acquire a set number of frames.  The user can also choose scaling type for a radiography type mode.  If neither scaling option is selected then no scaling is applied.  If both scaling options are selected, the resultant image is scaled up and down as shown in Figures 12b and 12c.

Furthermore, for a radiography type mode, the Use previous acquisition command option should not be selected for a Command Processor of revision H or later, which supports more options for manual and auto-sense X-rays selections.

The Calibration Setup group of settings allows the user to specify the number of frames used during calibration and the settings for the optional automatic offset calibration.

The Calibration Statistics group shows statistical information from the last gain calibration, which is stored on the host computer.  The Refresh button allows the user to retrieve the latest calibration statistical information from the command processor, display it and update the information stored on the host computer.

Figure 12a: Mode Settings dialog box for fluoroscopy type mode

Figure 12b: Mode Settings dialog box for radiography type mode with Use previous acquisition command option selected

Figure 12c: Mode Settings dialog box for radiography type mode with Use previous acquisition command option deselected

# V. Preferences

The Preferences dialog can be displayed via the Edit menu.  It allows for certain settings to be maintained between instances of running ViVA.  These settings are saved in the text file viva.ini, in the same directory as the application executable viva.exe.  The available settings are as follows:

- Client (PC) and server (Command Processor) IP addresses. Some validity checks are done on IP addresses, and the previous 7 IP addresses are kept in the list box for rapid reselection.

- Serial Port

- Connect by Ethernet or serial on starting ViVA

- Invert gray levels for fluoroscopy/radiography mode images and activate white area blanking

- Select white area blanking threshold and replacement values

- Select trilinear slope ratios

- Select window/level mapping function from linear, 2 non-linear or trilinear functions

- Save main window size, position and maximized state

- Use saved system settings such as Offset Corrections and Gain Corrections options from the ViVA’s text file (viva.ini) or use existing system settings in the Command Processor.

- Recently used file list that saves up to 16 entries of path and file names of the latest opened files.

The Preferences dialog box is shown in the following figure.

Figure 13: Preferences dialog box

# VI. Physics/Engineering Use

The Physics/Eng menu consists of commands that are intended for engineering purposes of a command processor.  These commands should be used with caution.

## VI-1. Display Test Pattern

If ViVA has an open connection to a VIP-9 command processor, choosing the Display Test Pattern option causes the command processor to display a test pattern.  This pattern consists of eight horizontal bars of equal height, with linear steps in intensity from zero at the top of the image, to 3584 counts at the bottom.  The image size is corresponds to the size of the image for the currently selected mode.

## VI-2. Set Debug Verbosity On/Off

The Set Debug Verbosity On/Off option causes the command processor to toggle between printing and not printing debug information over a serial port.

VI-3. Set Dcds Off

The Set DCDS Off command allows DCDS to be temporarily disabled which is useful in connection with performing analog offset calibrations. It will turn on again automatically when the mode is changed or offset calibration performed etc.

## VI-4. Transmit Files

The Transmit Files option allows user to transmit new files to the command processor.  Currently, ViVA can only transmit four types of files: (1) Receptor Configuration Data file, (2) Base Defect Map Image file. (3) Gain Calibrated Image file and (4) Offset Calibrated Image file.  An Ethernet connection to a VIP-9 command processor is required to use this command.  The Command Processor will immediately process this file and start operating using the new parameters contained in the file.

Figure 14: Transmit Files dialog box

## VI-5. Retrieve Files

The Retrieve Files option allows the user to retrieve a file from the command processor.  Currently, ViVA can only retrieve four types of files: (1) Receptor Configuration Data file, (2) Base Defect Map Image file. (3) Gain Calibrated Image file and (4) Offset Calibrated Image file.  An Ethernet connection to a VIP-9 command processor is required to use this command.  ViVA will store the newly retrieved file under the path and file name shown in the dialog box.

Figure 15: Retrieve Files dialog box

## VI-6. Analog Offset Calibration

The Analog Offset Calibration option allows the user to change analog offset settings and perform an analog offset calibration.  When selected, this option will display an Analog Offset Settings dialog as shown below.  Note that the Apply or OK button only updates the setting information at the command processor without performing a calibration.  The Calibration button initiates an analog offset calibration.

Figure 16: Analog Offset Settings dialog box

## VI-7. Image Arithmetic

The Image Arithmetic option allows the user to perform an arithmetic subtraction or addition on two images.  ViVA obtains a resultant image by performing a subtraction or an addition on a pixel of a source image and the corresponding pixel of a base image plus an offset factor.  These two images must be of the same size.

If the option Use Calculated Offset Factor is checked, ViVA will calculate this value by subtracting the two images and select the smallest negative integer of two pixels so that the minimum value of the pixels in the resultant image is zero after an addition of the calculated factor.  The absolute value of the integer will be used as the offset factor for the subsequent arithmetic option specified in the Image Arithmetic dialog box.  If the option is not checked, ViVA will use the offset factor number specified by the user in the text box on the right of the option.

Important!  Note that the file names in the forms “Subtracted-Image-n” or “Added-Image-n” where n is a positive integer, are used as temporary files during an image arithmetic operation.  Do not name or store any file in these forms under the directory where the executable viva.exe is located.

Figure 17: Image Arithmetic dialog box

## VI-8. Find Noisy Pixels

The purpose of the noisy pixel routine is to identify pixels that may not appear defective for each and every frame but over a number of frames fluctuate significantly in intensity.  The noisy pixel routine therefore collects a number of frames and attempt to identify noisy pixels statistically as those having a relatively high standard deviation (SD – also often referred to as “sigma”). The SD is calculated for each individual pixel (or binned grouping according to the mode); the SD is formed on a “temporal” not “spatial” basis as would be the case for example in ROI calculations.

Noisy pixel detection is begun by selecting this item from the Physics/Eng menu.  In the top section of the dialog box acquisition, parameters are selected.

Figure 18: Find Noisy Pixels dialog box

- Acquisition:

Each cycle will involve the acquisition of the specified number of frames and the retrieval of the last.  This allows the user to spread the actual frames used in the computation out over a larger number of frames, which could be useful for example if the fluctuations are clustered rather than random.

One frame is retrieved for each cycle; in ACCUMULATION modes this is actually the sum of 2 frames formed in the command processor. In order to compute the SD effectively, it is suggested that at least 10 cycles be specified, but, for initial familiarization purposes, 5 will be adequate.  The acquisition may take some time particularly in modes where no binning is occurring and will be essentially proportional to the number of cycles.  The user may choose to save each individual retrieved frame to a disk file via the check box; this feature will probably not be selected normally but may be found useful on occasion.

Acquisition is begun by clicking the Acquire Noisy Pixel Data button.  At the end the user is given the opportunity to save to disk the SD image formed. It will probably be found convenient to do this, at least saving the data temporarily.  The SD image is actually the calculated standard deviation scaled by 10 to improve precision.  Finally a thumb view of the SD image is displayed.

- Set Threshold and Display:

When acquisition is complete a threshold can be selected above which pixels are considered defective.  Changing the value of the threshold will normally result in a corresponding change in the indicated number of pixels above the threshold.  The displayed number is updated whenever the threshold is changed and focus shifted away from the set threshold edit box (e.g. when finished typing press “TAB” key).

When the required threshold has been selected, a defect map can be created by pressing the Display Noisy Pixels button.  Note that if this button is clicked when no acquisition is current (i.e. no acquisition done since dialog box opened), the user is given an opportunity to select a previously acquired SD image from a disk file.

The choice of the threshold is important in achieving the goal of accurately identifying truly noisy pixels.  One method to do this requires exiting the dialog box and inspecting the acquired SD image.  Magnify some part of it and look for a featureless – “normal” – region.  Use the ROI facility to evaluate the average SD and the standard deviation of the SD itself in this region, and repeat in other areas to get good numbers for the average normal SD and its distribution as indicated by the standard deviation box.  A good choice for the threshold will likely be 6x the standard deviation plus the average. This procedure is what is carried out automatically by clicking the AutoSet button.

- Updating the base defect map:

When the required threshold is identified the Noisy Pixel dialog can be reopened if necessary and the data of interest reactivated by pressing the Display Noisy Pixels button, and selecting the appropriate file.  Set the threshold and click the Display Noisy Pixels button again to ensure the defect map is created at the required threshold setting.  When sure that this map is correct, download to the command processor by clicking the OR With Aux Defect Map or OR With Base Defect Map button.  This button results in the logical OR operation being performed between the existing base defect map and the created noisy pixel defect map. If the threshold value is changed without creating a new map, then a new map is automatically created prior to updating the base defect map.  The result is downloaded to the command processor.

## VI-8. Defect Map Editor

The Defect Map Editor is invoked by choosing the option Physics/Eng|Defect Map Editor.  This will switch ViVA from image viewing mode to image editing mode.  In this mode, the Acquisition toolbar is replaced by the Defect Map Editor toolbar as shown below.  The Status Bar is displayed to show the currently selected mode in the Command Processor.  All images in the Image View are closed and the Image Layout is set to Two-Image Layout.  Then ViVA retrieves the Base Defect Map image and the Auxiliary Defect Map image from the Command Processor and displays them in the Thumbnail View and the Image View.  The Base Defect Map image is placed in the right pane and the Auxiliary Defect Map image is placed in the left pane.

Figure 19: Defect Map Editor toolbar

Figure 20: Defect Map Editor screen shot

The editing mode is only activated by the selection of the Standard Selection Cursor.  All other cursors work in Defect Map Editor mode as they do in the viewing mode.  Thus, to edit a point or a segment, the user must click on the Standard Selection cursor first.  ViVA displays good pixels in black and defective pixel in white.

### VI-8-1. Editing A Pixel

On the leftmost of the Defect Map Editor toolbar is the pixel list, which contains all edited pixels as they are marked from good to defective.  Please, note that this list contains only the manually edited pixels.  Defective pixels found on the retrieved or open image are not included in the list.

- With mouse actions: To mark a good pixel as defective or a defective pixel as good, simply point the cursor at the pixel and click on the left button.  When a good pixel is marked as defective, the pixel is placed on the pixel list and when a defective pixel is marked as good, the pixel is deleted from the pixel list if it exists in the list.

- With keyboard actions: Similar to the mouse actions, the arrow keys and the Enter key can be used to edit a pixel.  Use the arrow keys or the mouse to point a desired pixel.  Then press the Enter key to toggle the pixel from good to defective or defective to good.

- With the pixel list:  (1) To add a pixel onto the pixel list, select the entry < Pixel >, type the coordinates of the pixel in the point form (x, y) and then press the Tab key to end the addition.  (2) To delete a pixel, select the desired pixel, press the Backspace key and then press the Tab key to end the deletion.  (3) To change the coordinates of a pixel, select the pixel, edit the x and y coordinates and then press Tab key to end the change.

### VI-8-2. Editing A Row Segment

On the second leftmost of the Defect Map Editor toolbar is the row segment list, which contains all edited row segment as they are marked from good to a defective.  Please, note that this list contains only the manually edited row segment.  Defective row segments found on the retrieved or open image are not included in the list.

- With mouse actions: To mark a good row segment as defective or a defective row segment as good, simply point the cursor at the row and then click on the left button while holding down the Shift key.  When a row segment is marked as defective, the segment is placed on the row segment list and when a defective row segment is marked as good; the segment is deleted from the row segment list if it exists in the list.

- With keyboard actions: Similar to the mouse actions, the arrow keys and the Enter key can be used to edit a row.  Use the arrow keys or the mouse to point a desired row.  Then press the Enter key while holding down the Shift key to toggle the whole row from good to defective or defective to good.

- With the row segment list:  (1) To add a row segment onto the row segment list, select the entry < Row Segment >, type in the coordinates of the row segment in the row segment form (x1, y) to (x2, y) and then press the Tab key to end the addition.  (2) To delete a row segment, select the desired row segment, press the Backspace key and then press the Tab key to end the deletion.  (3) To change the coordinates of a row segment, select the row segment, edit the x and y coordinates and then press Tab key to end the change.

### VI-8-3. Editing a Column Segment

On the third leftmost of the Defect Map Editor toolbar is the column segment list, which contains all edited column segment as they are marked from good to a defective.  Please, note that this list contains only the manually edited column segment.  Defective column segments found on the retrieved or open image are not included in the list.

- With mouse actions: To mark a good column segment as defective or a defective column segment as good, simply point the cursor at the column and then click on the left button while holding down the Control key.  When a column segment is marked as defective, the segment is placed on the column segment list and when a defective column segment is marked as good; the segment is deleted from the column segment list if it exists in the list.

- With keyboard actions: Similar to the mouse actions, the arrow keys and the Enter key can be used to edit a row.  Use the arrow keys or the mouse to point a desired row.  Then press the Enter key while holding down the Control key to toggle the whole column from good to defective or defective to good.

- With the column segment list:  (1) To add a column segment onto the column segment list, select the entry < Column Segment >, type in the coordinates of the column segment in the column segment form (x, y1) to (x, y2) and then press the Tab key to end the addition.  (2) To delete a column segment, select the desired column segment, press the Backspace key and then press the Tab key to end the deletion.  (3) To change the coordinates of a column segment, select the column segment, edit the x and y coordinates and then press Tab key to end the change.

### VI-8-4. OR Image Operation

To add all defective pixels on the Auxiliary Defect Map on the left Image View pane to the Base Defect Map image on the right of the Image View pane, click the OR Image button.  This command is useful when there is an existing image with defective pixels that can be added onto the Base Defect Map image.  The user simply opens an image as he/she normally would; then drag and drop the thumb view of the image onto the left pane before OR-ing the two images.

### VI-8-5. Retrieve Images

To retrieve a new pair of a Base Defect Map image and an Auxiliary Defect Map image, click on the Retrieve button.  This command allows the user to restart editing images without exiting the Defect Map Editor mode.

### VI-8-6. Transmit Image

To transmit the Base Defect Map image to the Command Processor, click on the Transmit button.  This transmits the image on the right Image View pane to he Command Processor and overwrites the current Base Defect Map Image with the new one.

### VI-8-7. Close Defect Map Editor

To close switch back to ViVA viewing mode, click on the Close button on the Defect Map Editor toolbar.  All images in the Thumb View and the Image View remain unchanged.  The Defect Map Editor toolbar is replaced by the Acquisition toolbar.

# VII. Statistics

The Statistics menu consists of two commands that allows the user to obtain statistical information on a specific region of an image called the Region Of Interest (ROI).  There are two commands available under the Statistics menu, which are described as follows:

- The Select Whole Image option allows the user to select the whole active image as a region of interest.

- The Show ROI Statistics… option allows the user to display statistical information.  It also allows the user to redefine the region of interest by entering two new diagonal points of a new region.  There are two command buttons in this dialog box.  The Apply button causes recalculation of all statistical information using the current two points specifying a region and the Close button closes the dialog box.  The statistical information provided on this dialog box is explained in the following table:

Table 2: Descriptions of ROI statistical information.

| Name | Format | Description |
| --- | --- | --- |
| 01. Region of Interest (ROI) | (x0, y0) to (x1, y1) | Two diagonal points that define the ROI. |
| 02. Width and Height of Region | integer x integer | Width and height dimensions of the ROI. |
| 03. Center of ROI | (x, y) | Center location of the ROI. |
| 04. Number of Pixels in ROI | integer | Number of pixels found in the ROI. |
| 05. Number of Defective Pixels in ROI | integer | Number of defective pixels found in the ROI. |
| 06. Percent of ROI/Image | real | Percent of the region over the whole image. |
| 07. Maximum Pixel Value in ROI | integer @ (x, y) | Maximum pixel value first found in the ROI. |
| 08. Minimum Pixel Value in ROI | integer @ (x, y) | Minimum pixel value first found in the ROI. |
| 09. Total of Pixel Values in ROI | integer | Total of all pixel values found in the ROI. |
| 10. Average of Pixel Values in ROI | real | Average of pixel values found in the ROI. |
| 11. Standard Deviation in ROI | real | Standard deviation of pixel values in the ROI. |
| 12. Signal/Noise Ratio in ROI | real | Defined by formula average/standard deviation. |

The coordinates (x, y) of a point and the dimensions width by height are limited by the size of an image.  Specifically, ranges of a pair (x, y) are from 0 to (width – 1) and 0 to (height – 1), respectively.  The two image sizes currently used are 768 x 960 and 1536 x 1920.  Pixel values are in range from 0 to 4095.  The pixel value of 1 indicates a defective pixel.

ViVA also provides other methods to specify a region of interest as shortcuts.  The user can use the left mouse button and drag techniques to draw a rectangle on an image to indicate a region of interest.  Then the user can use the command Show ROI Statistics… under the Statistics menu or the radio button ROI Stats… on the vertical tool bar to display the statistical information.  Similarly, the user can immediately obtain the information after drawing a region by holding down a Ctrl key while dragging the mouse.  Note that the Ctrl key must be pressed and held down before dragging is used to specify a region of interest.

Figure 21: Region of Interest Statistics dialog box

Note that the Text option in the dialog box above allows the user to save the pixel values of the selected ROI as text values in a file, which can be can read by a word processor such as Microsoft Word or a spreadsheet program such as Microsoft Excel.  These text values are separated by a tab key and a row is ended with an Enter key.