# Table of Contents

1.	Introduction	2

1.1	Purpose	2

1.2	Scope	2

1.3	Definitions, Acronyms, and Abbreviations	2

2.	Assembly Instructions	3

3.	Release Notes	4

3.1	Software Build/Release	4

3.2	Changes from Previous Revision	4

3.3	Hardware Compatibility	11

3.4	Upgrade Considerations	11

3.4.1	Manufacturing	11

3.4.2	Upgrade Installed Systems from Software Rev. A, B, or C	11

3.4.3	Upgrade Pre-Rev. A Installed Systems	11

3.4.4	Revision Labeling	12

4.	Software and Data Components	13

4.1	CD Index	13

4.2	System- and Receptor-specific Data	13

# Introduction

## Purpose

This document describes the contents of the compact disc (CD) which contains the released software and data files for the PaxScan family of Imager systems. This software and these data files together constitute the PaxScan “Software Package”.

## Scope

Included in the Software Package are all executable programs and data which are:

- Developed by VIP product engineering staff, and

- Required for the manufacturing, setup, and testing of a PaxScan system, and

- Are needed by OEM customers for interfacing with a PaxScan system.

This document also describes some of the data, which is produced in the manufacturing of a PaxScan system and which must be sent to the OEM customer and archived by manufacturing and by VIP.

## Definitions, Acronyms, and Abbreviations

This section provides the definitions of all terms, acronyms, and abbreviations required to properly interpret this document.

CD	Compact Disc

CCN	Change Control Number; also used to describe an Engineering Change Order

VIP	Varian Imaging Products

# Assembly Instructions

This section describes how the Software Package is assembled. The parts required to assemble the Software Package are shown in Table 1, below.

Table 1: Software Package Components

| Component | Material Number |
| --- | --- |
| Software Package (top level) | 22679 |
| Blank CD-R Disc | 7850093400 |

For a new revision, the Software Engineering group is responsible for providing the correct components listed in the CD Index (see Section 4.1 of this document).   For further instructions on how to build the Shipping CD, see the PaxScan 2520 Configuration and Test Procedure (42103001_003).

A Software Configuration Management tool is used to ensure that the correct version of each component is selected. Software Engineering will copy these components onto the blank CD-R disc, and will then label the disc using the CD Label. The CD Label will include the following information:

Title:	Software Package

Part #:	22679

Date:	3/14/2005

Rev.:	J.05, Bld 66

Software Engineering will then complete a CCN. The complete release consists of a hard copy of the Software Package Description (this document), the completed, labeled CD, and the accompanying CCN.

# Release Notes

## Software Build/Release

All software and data components are controlled and archived using a Software Configuration Management system. Regular builds are performed by the Software Engineering group. Each build is numbered consecutively. The revision letter is incremented for a new software release.

For purposes of configuration management in the Software Engineering group, this Revision is designated as follows:

- Revision J.05

- Build 66

Accordingly, in the Software Configuration Management system, all components of this build have been labeled as follows:

- BLD_J05.066

## Release Package Sub-Component Version Numbers

| Component Name | Component Location | Version Number |
| --- | --- | --- |
| Configure.exe | Release_dir\Configure\ | Rev J.05 Build 66 |
| Viva.exe | Release_dir\Viva\ | Revision J.05 Build 66 |
| Hostdown.exe | Release_dir\Command_processor\ | Version 1.0 |
| System_Software.bin | Release_dir\Command_processor\ | Rev J.05 Build 66 |
| Global_Control_Firmware | Release_dir\Command_processor\ | GC Rev J, Build 06 |
| Dcds_Receptor_Firmware | Release_dir\Command_processor\ | DCDS Rev J, Ver 06 |
| Std_Receptor_Firmware | Release_dir\Command_processor\ | STD Rev J, Ver 01 |
| 4.1_dcds_receptor_firmware | Release_dir\Command_processor\ | 4.1 Recpt Firm ver033104 |

## Changes from Previous Revision

Changes that are new with this revision are shown here. For a complete change history, see Appendix A.

Table 2 Changes from J.05 Build 65

| Changes in Rev J.05 Bld 66 | Affected Files | ClearQuest Record # |
| --- | --- | --- |
| Modified 4.1 PS2520 Receptor firmware to work with and new Gate Driver, TMS57600A. | 4.1_dcds_receptor_firmware.hex | 0862 |

Table 2: Changes to rev J.05

| Changes in Rev J.05 | Affected Files | ClearQuest Record # |
| --- | --- | --- |
| Added support for Paxscan 4.1 2520 DCDS receptors | Configure.exe 4.1_dcds_receptor_firmware.hex | 612 651 |

## Table 3: Changes to rev J.04

| Changes in Rev. J.04 | Affected Files | ClearQuest Record # |
| --- | --- | --- |
| Fixed a problem with new 2520 power supplies (420163-01) blowing fuses. It has been determined that the voltage settings need to change, along with a ECO modification to the power supply hardware, to fix the problem | Configure.exe | 541 |
| Added Fiber-Optic Power supply to System Configuration window | Configure.exe | 569 |
| Panel Ready should remain low (de-asserted) until the system has completed boot-up. Once the system has boot, panel ready should go high alerting the user that the system is now ready to accept valid xray. | System_software.bin | 530 |
| In the Configure Application it was possible to set two rad modes or 2 full res modes, an illegal configuration. This has now been fixed, allowing only one rad or full res. mode. | Configure.exe | 529 |
| Under certain conditions, flash files could be added to the flash that are never deleted, particularly when the system software is loaded in the wrong slot. This would cause the flash to feel up with files that could not be erased. | Vip_proms.hex | 519 |
| Problem retrieving defect and Aux Defect maps from 2520 full res modes, or 4030a full res modes would fail with some new motherboards was fixed in the vip_comm.dll | Vip_comm..dll | 509 |
| The configure program will allow a user to set 4 frames per second for full res. low noise fluoro, but VIVA does not recognize it and reports a blank when opening the "Mode Settings". Problem was determined to be 4 fps is not support by command processor hardware, and was removed from the frame rate selection window in configure. | Configure.exe | 420 |
| Fixed a problem in Viva noisy pixel defects not been able to be displayed after acquiring the noisy pixel data. | Viva.exe | 572 |
| Add support for new 2520 power supplies (420163-01). | Configure.exe | 396 |
| With J.03 the recursive filter settings were changed to not be applied until a specified number of frames passed. This was added to eliminated any image defects caused by the mode change. In previous versions any defects would be picked up by the recursive filter, resulting in the defects persisting in the image. However, while J.03 fixed this problem, it had the side effect of only updating the recursive filter at the assertion of prepare. Older versions would allow the recursive filter to be changed while the mode is running. This was fixed in this version (J.04) system software. | System_software.bin | 535 |
| Added support for VMS Mac Addresses (OUI Block of 00:07:02). | System_software.bin VxWorks.hex | 531 |

Table 4: Changes to rev J.03

| Changes in Rev J.03 | Affected Files | ClearQuest Record # |
| --- | --- | --- |
| Change video timing for 420080-[01 and 02] receptors to match the fastscan modes that are available for 420080-[05 and 06] receptors. | Configure.exe | 294 |
| Fixed problem with Gain of 8 receptors (420080-07) reporting gain of 4 instead of gain of 8. | System_software.bin | 339 |
| Modify status text content in Viva to clarify usage for radiography gain calibration and image acquiring. | Viva.exe | 326 |
| Changed find-noisy-pixels test in Viva to automatically turn off all corrections before acquiring data. | Viva.exe | 256 |
| Fixed a problem with first few frames in full res. Fluoro being corrupted when recursive filter is enabled. | System_software.bin | 323 |
| Added ability to perform radiography gain calibrations using Autosense with N-Frames. | System_software.bin, Viva.exe | 292 |
| Fixed a problem with command processor that have an ABS board installed,not booting correctly when a low-noise-radiography or low-noise-full-res. fluoro mode is loaded and the frame rate set to 2fps | System Software.bin | 353 |
| Added Defect Analysis to defect map editor in ViVA. Before transmitting defect maps to the command processor, the defect map editor will check newly created defect map to ensure receptor defects are within published specification. | Viva.exe | 35 |
| Added support for gain of 8 receptor type (420080-07) | System_Software.bin, Configure.exe | 298 |

Table 5: Changes to rev J.02

| Changes in Rev J.02 | Affected Files | ClearQuest Record # |
| --- | --- | --- |
| Fixed problem with Analog Offset calibrations not working correctly. Use a median percent value of 1 to utilize improvements in auto offset algorithm. | System_software.bin | 205 |
| Fixed Problem with assertion of Prepare or Valid-Xrays during boot-up sequence crashing command processor. | System_software.bin | 260 |
| Fixed error in sample code that caused radiography acquisitions with more than 1 frame to be aborted prematurely. | System_software.bin | 285 |
| Fixed bug with get vip_get_system_info was returning Max columns and Max lines wrong. In previous versions the values were reversed. | System_software.bin | 284 |
| Repaired problem with Auto-sense feature not working correctly after canceling offset calibration. | System_software.bin | 282 |
| Fixed Problem a with retrieving images after an offset calibration. This problem would only occur when the user acquires an image than waits some amount of time before retrieving the image. If an auto offset calibration was initiated during the wait, the image retrieved would be incorrect. | System_software.bin | 281 |
| Fixed problem with Viva not reporting errors correctly when retrieving images. | Viva.exe | 275 |
| Line Noise rejection is now always disabled in Viva | Viva.exe | 274 |
| Fixed several bugs with auto-offset calibrations not working correctly | System_software.bin | 265/264 |

Table 6: Changes to rev J.01

| Changes in Rev J.01 | Affected Files | ClearQuest Record # |
| --- | --- | --- |
| Repaired bug that caused the command processor to crash when canceling a OEM generated offset calibrations, as is the case when using ViVA. | System_software.bin | 202 |
| Repaired bug with the first frame of the first pulse being missed in pulsed gain calibrations. | System_software.bin |  |
| Fixed bug with zoom mode’s binning not being displayed correctly when using standard panels. | System_software.bin |  |

Table 7: Changes to revision J

| Changes in Rev J | Affected Files | ClearQuest Record # |
| --- | --- | --- |
| Ability to turn off AGC from ViVA. | ViVA.exe | 18 |
| Autosense values zero added to error checking. | system_software.bin | 27 |
| DCDS analog offset added. | system_software.bin | 32 |
| DCDS config data updated. | configure.exe | 33 |
| Ability to prefix version header to receptor and GC firmware was added configure.exe. | configure.exe system_software.bin | 34 |
| .viv file format added to ViVA. | ViVA.exe | 55 |
| Make mode changing in state machine smarter. | system_software.bin | 67 |
| Manufacturing GUI doesn't match draft sheet. | configure.exe | 68 |
| Provide source code samples. | communications_interface_design.doc | 73 |
| Move non-public commands to non-public area in comm_dll. | vip_comm.dll | 86 |
| Implement get_system_version_numbers command. | vip_comm.dll system_software.bin | 89 |
| Complete firmware mods for DCDS. | global_control_firmware.hex dcds_receptor_firmware.hex | 90 |
| ViVA ROI not working correctly. | ViVA.exe | 99 |
| Rad scaling broken. | system_software.bin | 120 |
| VIVA: Rad Mode Gain Calibration GUIs incorrect. | ViVA.exe | 143 |
| Flash doesn't handle 3 flouro modes. | system_software.bin | 150 |
| Add Rad Scaling, # cal frames and auto cal to manufacturing configure interface | configure.exe system_software.bin | 151 |
| Expose OK not working. | global_control_firmware.hex | 152 |
| ViVA: Delay during defect map download process added. | ViVA.exe | 162 |
| Can't retrieve test image with ViVA. | ViVA.exe | 166 |
| Needed a convenient way to turn off DCDS. | ViVA.exe vip_comm.dll system_software.bin | 167 |
| Print IP address to serial output at startup. | system_software.bin | 171 |
| Current image request needs verification. | system_software.bin | 172 |
| Better identifying mechanism for zoom modes. | system_software.bin | 173 |
| Add mode acq type to config app. | configure.exe | 175 |
| Incorporate console configure into GUI based config_app for backward compatibility. | configure.exe | 176 |
| Add IP address list and check IP address in ViVA. | ViVA.exe | 178 |
| Repair path/file names in save dlg box in ViVA. | ViVA.exe | 179 |
| Disable BLC. | system_software.bin | 182 |

# Hardware Compatibility

The revision level for each piece of hardware to be compatible with this revision software is given below, along with other relevant information.

- This revision no longer supports first generation ARM prototype Receptors.

- This revision is not compatible with System Configuration Files and Receptor Configuration Files used in pre-release versions of the software, e.g. “Release 0” and “Release 1.” In subsequent revisions, all efforts will be made to maintain backward compatibility of system and receptor configuration information, in order to simplify software upgrades of existing units.

- The required revision level for the Command Processor motherboard is Rev. D or later. Gain/offset buffer memory configuration 32 MB only is supported.

- The required revision level of the Command Processor Global Control board is Rev. E or later.

- The required revision level for the (optional) ABS/Video board is Rev. 1.2C or later.

## Upgrade Considerations

### Manufacturing

- Manufacturing should change over to using all the new software and firmware provided with this revision.

### Upgrade Installed Systems from Software Rev. A, B, or C

- Hardware should be updated to meet the requirements listed in Section 3.3, “Hardware Compatibility.”

- Use hostdown.exe to download system_software.bin, global_control_firmware.hex, std_receptor_firmware.hex (or dcds_receptor_firmware.hex as appropriate) and (if the ABS/Video Board is installed) abs_firmware.bin to the Command Processor. Also, the appropriate configuration files must be downloaded via hostdown.exe.  Reboot the Command Processor and wait 2 minutes. Reboot the Command Processor once more, and it will be ready to use.

- A new version of ViVA is available by copying the files found in the directory called “ViVa” to hard disk and running “ViVa.exe.

### Upgrade Pre-Rev. A Installed Systems

- Hardware should be updated to meet the requirements listed in Section 3.3, “Hardware Compatibility.”

- The format and content of system and receptor configuration data files in the current revision is not compatible with pre-Rev. A versions. Please contact Technical Support at  (1-800-INTERAY) for assistance with updating these files (typically, system_config.dat and receptor_config.dat.)

### Revision Labeling

- A Letter change denotes a modification has been made to one or more of the VIP commands used for communication between the command processor and the host computer.

- The Numerical number denotes bug fixes made to the software release with no modifications to VIP commands.

- Any software component version of a letter reversion level can communicate to any other software component with the same letter version.

- When upgrading a software component to a new letter revision, all the other software components must be upgraded as well.

# Software and Data Components

The Software Package includes software and data files which are required by manufacturing and by OEM customers. There are also Receptor-specific data files which should be shipped to the OEM customer with each VIP Receptor.

All components of the Software Package will be supplied to manufacturing and Imaging Products engineering.

For each new receptor, manufacturing should create a “Receptor CD,” which includes some of the contents of the Software Package CD, and also the receptor-specific data (see Section 4.1, following). One copy should go with the receptor, one copy should go to Imaging Products engineering, and one copy should be kept at the manufacturing site. The CD’s shall be built in accordance with the instructions specified in the Paxscan 2520 Customer Configuration and Test Procedure  (42103001_003).

Imaging Products engineering will create an additional archive for external release, to support research and development, continuing engineering, and field service. This archive will be placed in a “zipped” file, “software_package.zip”, on the Software Package CD.

## CD Index

The contents of the Software Package CD are described in the System Config and Test Document (42103001_003).

## System- and Receptor-specific Data

For each receptor that is manufactured, there are configuration data and defect map information that is specific to that unit. The information that is compiled during manufacturing/setup/test is shown in Table 3, below.

Table 3: System- and Receptor-specific Data

| Component | Suggested File Name | Receptor CD | Comment |
| --- | --- | --- | --- |
| System Configuration File (Binary) | system_config.dat | Yes | Binary version is contained in Command Processor. |
| System Configuration File (Text) | system_config.txt | Yes | Human-readable version. |
| Receptor Configuration File (Binary) | receptor_config.dat | Yes | Binary version is contained in Command Processor. |
| Receptor Configuration File (Text) | receptor_config.txt | Yes | Human-readable version. |
| Defect Maps | Defect_Map_0.raw, …1, …2, … | Yes | Create during mfg./setup/test for each receptor, one for each mode; names correspond to mode numbers, 0, 1, 2, etc. |

Appendix A: Change History

Rev. A

- First release.

Rev. B

- Fixed problem with bad data in first frame of Rad mode acquisition.

- Application no longer waits for a Frame Start event when turning the IPU off (no user impact).

- A 20 second delay was added between the flat field and dark field acquisition portions of gain calibration. This will ensure that any image lag in the receptor does not invalidate the dark field acquisition.

- File names were changed from “vip_ethernet.*” to “vip_comm.*”, in anticipation of adding serial communications capability to the current ethernet-only connection. This will require that any client application which uses the communications DLL be recompiled.

- Added capability to change frame rate in receptor configuration file.

- Added capability to change saturation threshold settings in receptor configuration file.

- Updated Global Control board firmware in anticipation of future feature enhancements.

- CreateDefectMap and ViVA applications recompiled to use new communications DLL names.

Rev. C

- Applied fix for continuing problem with bad data in first frame of Rad mode acquisition.

- Changes to Rad mode acquisition to increase robustness (resistance to crashing).

- Post-acquisition processing for Rad mode has been modified. In Rad mode, the system currently integrates the results of a set number of frames. Previously, the resulting accumulated image was normalized by dividing by the number of frames accumulated, before displaying the result. Also, there was a bug in the algorithm, which had the effect of reducing the displayed output by an additional factor of two. In this new revision, the accumulated image is displayed “as is,” unless so much dose has accumulated that the resulting image would not “fit” in the required 12-bit output range. The determination of excess dose is made based on a histogram of pixel values from the central one-ninth of the image. The requirement is that 80% of those pixels should have values less than one-half of the system’s pixel saturation threshold (which is user-definable). If this requirement is not met, then a single linear scaling value is calculated and applied to all pixels in the image, such that the requirement is satisfied. The value of saturated pixels is set to the system’s saturation replacement value (also user-definable).

Rev. D

- Completely new version of ViVA based on Visual C/C++ rather than Visual Basic.

- ABS firmware updated to support the following video modes: 525 lines / 60 Hz, 1049 lines / 60 Hz, 625 lines / 50 Hz, and 1249 lines / 50 Hz. Command Processor software and Configure application updated to support this change.

- New versions of Global Control and Receptor firmware to support auto-sense rad operation and zoom modes.

- Support for 3 zoom modes added: Zoom Center, Zoom Center Left, Zoom Center Right.

- Communications library updated to eliminate need to call the function vip_register_application. See vip_comm.h for interface changes. Backward compatibility with old calling style maintained.

- Support for serial interface included.

Internal bug fixes and changes for Command Processor application…

- Previously, when changing modes, only changed register values were written to the Global Control board; this was not working correctly; now all values are written.

- Eliminated ARM receptor code.

- Rad mode was hanging up; changed so IPU is set up before the Global Control board.

- Front panel LED operation enabled.

- Changes to support auto-sense radiography operation (not yet released.)

- Previously, the IPU would hang upon IPU on/off if not done on a frame end. The problem was resolved in hardware for Rev. D motherboards, so the software workaround was removed.

- In a system having 3 modes, and then used only for a single rad mode, the flash files for the “extra” modes were not deleted. code added to delete the extra modes.

- Fixed bug in command processor that required a reboot after using the interface function vip_put_config_data(…).

Rev. E

- New versions of Global Control and Receptor firmware to support timing discrepancies as found in Rev D.

- The DOS based configure application used to create the files, system_config.dat and receptor_config.dat has been replaced with a GUI interface.

- Auto-sense radiography mode available and radiography gain calibration changed to auto-sense radiation as well as to accumulate multiple radiation shots.  The communications command gain_cal_prepare added a parameter to indicate auto-sensed gain calibration.  See the document vip_comm/communications_interface_design.doc for details.

- Configuration files in flash and any files opened with Configure.exe are automatically updated to include a new token scheme so files can be backwards compatible.  Once a config file is opened using the Rev E version of configure.exe, that file cannot be opened with older versions of configure.exe.

- Analog offset calibration was added to automatically update hardware offset values.  This included adding new communications commands (analog_offset(), get_analog_offset_data(), set_analog_offset_data()) as well as adding new menu items to ViVA.  See the documents vip_comm/communications_interface_design.doc and viva/ViVA_Help.doc for more details.  Configure.exe was also changed to set the required data to perform the analog offset.

- Gain calibration statistics were added to the gain calibration calculation and are available for display using ViVA or the communications command, get_cal_stats().See the documents vip_comm/communications_interface_design.doc and viva/ViVA_Help.doc for more details.

- Radiography scaling choice was added so that the image is not scaled or scaled up/down/both. This included adding new communications commands (set_rad_scaling(), get_rad_scaling ()) as well as adding new menu items to ViVA.  See the documents vip_comm/communications_interface_design.doc and viva/ViVA_Help.doc for more details.  Configure.exe was also changed to set the required data.

- Added startup behavior choices to be: active acquisition, display test pattern and wait for command or stand-alone.  This choice must be specified in the system config GUI of configure.exe and the command processor software will react accordingly.  The communications command get_system_info was changed to added the startup configuration parameter, but is backwards compatible. See the document vip_comm/communications_interface_design.doc for more details.

- Added menu item to ViVA to allow enabling/disabling debug print statements over serial. See the document viva/ViVA_Help.doc for more details.

- Added parameters ready for pulse and number of pulses to the communications command query_progress, but the command is backwards compatible. See the document vip_comm/communications_interface_design.doc for more details.

- Changed receptor config data defaults for radiography to be frame rate of 1 frame/sec, 8 frames for offset and gain of 1.

Rev. F

- Added feature to subtract two files in ViVA.

- Added ability to retrieve the receptor config file from the command processor via ViVA.

- Added capability to provide statistics for a region of interest in ViVA.

- When choose serial port preferences, ethernet preferences are disabled.

- Banding problem with Radiography mode eliminated.

- Configure.exe default values updated to match Rev E firmware changes.

- Canceling gain calibration fixed so system does not hang.

Rev. G

- Fixed External Sync. and Line Sync modes in Receptor Firmware.

- Changed Configure default parameters for Zoom Mode.

- Modified System Software build label.

- Modified Receptor Firmware to eliminate vertical line artifacts on readout ASIC boundaries.

- Modified Receptor Firmware to correct the addressing of the register, which programs the analog “TFT-ON” voltage, AVGH.

Rev. H

- Added DCDS capability.

- Added new DCDS modes.

- Added a new panel type called DCDS.

- In all acquisition-type functionality, reversed the usage of frame start and frame end since the registers get latched after frame start instead of frame end.

- In all accumulation-type functionality, the number of frames delayed before ready for acquisition was increased to 10 so the hardware will be ready.

- Error checking added to prevent possible mathematical impossibilities.

- For analog offset calibration, the ability to turn DCDS on/off was added.

- Added new communications commands called vip_get_mode_acq_type and vip_set_mode_acq_type.

- Repaired the problem that the ethernet interface is disabled if an auto offset is cancelled.

- Repaired the problem that the BLC correction is TRUE even if the BLC count is zero.

- Repaired the problem that if radiography was mode zero, could not perform a gain calibration for radiography.

- Added Graphical version of defect map editor.

- Automatic noisy pixel detection (post-processing).

- Rad valid frame signal was added for accuracy.

- Added command vip_set_frame_rate to change the frame rate.

- Modified System Software build label.

- Repaired problem with analog offset not working correctly with zoom modes.

- Repaired problem with command processor and ViVA becoming out of sync after offset or analog offset calibrations.

- Improved window and leveling algorithm in ViVA.

- Added new hardware signal to external Sync Port (Expose_OK  pins 9 and 10).  This signal indicates when it is safe to expose the receptor to a pulse of radiation.

- Added new hardware signal to external Sync Port (ValidRadFrame  pins 11 and 12).  This signal indicates when a fully processed rad frame is available on the 16 bit video port.

Rev. J

- Ability to turn off AGC from ViVA.

- Autosense values zero added to error checking.

- DCDS analog offset added.

- DCDS config data updated.

- Ability to prefix version header to receptor and GC firmware was added configure.exe.

- .viv file format added to ViVA.

- Make mode changing in state machine smarter.

- Manufacturing GUI doesn't match draft sheet.

- Provide source code samples.

- Move non-public commands to non-public area in comm_dll.

- Implement get_system_version_numbers command.

- Complete firmware mods for DCDS.

- ViVA ROI not working correctly.

- Rad scaling broken.

- VIVA: Rad Mode Gain Calibration GUIs incorrect.

- Flash doesn't handle 3 flouro modes.

- Add Rad Scaling, # cal frames and auto cal to manufacturing configure interface.

- Expose OK not working.

- ViVA: Delay during defect map download process added.

- Can't retrieve test image with ViVA.

- Needed a convenient way to turn off DCDS.

- Print IP address to serial output at startup.

- Current image request needs verification.

- Better identifying mechanism for zoom modes.

- Add mode acq type to config app.

- Incorporate console configure into GUI based config_app for backward compatibility.

- Disable BLC.

Rev. J01

- Repaired bug that caused the command processor to crash when canceling a OEM generated offset calibrations, as is the case when using ViVA.

- Repaired bug with the first frame of the first pulse being missed in pulsed gain calibrations.

- Fixed bug with zoom mode’s binning not being displayed correctly when using standard panels.

Rev. J02

- Fixed problem with Analog Offset calibrations not working correctly.  Use a median percent value of 1 to utilize improvements in auto offset algorithm.

- Fixed Problem with assertion of Prepare or Valid-Xrays during boot-up sequence crashing command processor.

- Fixed error in sample code that caused radiography acquisitions with more than 1 frame to be aborted prematurely.

- Fixed bug with get vip_get_system_info was returning Max columns and Max lines wrong.  In previous versions the values were reversed.

- Fixed error in sample code that caused radiography acquisitions with more than 1 frame to be aborted prematurely.

- Line Noise rejection is now always disabled in Viva.

- Fixed several bugs with auto-offset calibrations not working correctly.

Rev. J03

- Change video timing for 420080-[01 and 02] receptors to match the fastscan modes that are available for 420080-[05 and 06] receptors.

- Fixed problem with Gain of 8 receptors (420080-07) reporting gain of 4 instead of gain of 8.

- Modify status text content in Viva to clarify usage for radiography gain calibration and image acquiring.

- Changed find noisy pixels test in Viva to automatically turn off all corrections before acquiring data.

- Fixed a problem with first few frames in full res. Fluoro being corrupted when recursive filter is enabled.

- Added ability to perform radiography gain calibrations using Autosense with N-Frames.

- Fixed a problem with command processor that have an ABS board installed, not booting correctly when a low-noise-radiography or low-noise-full-res. fluoro mode is loaded with the frame rate set to 2fps.

- Added Defect Analysis to defect map editor in ViVA.  Before transmitting defect maps to the command processor, the defect map editor will check newly created defect map to ensure receptor defects are within published specification.

- Added support for gain of 8 receptor type (420080-07)

Rev. J04

- Fixed a problem with new 2520 power supplies (420163-01) blowing fuses.

- Panel Ready should remain low (de-asserted) until the system has completed boot-up.

- Fixed problem with Configure Application allowing two rad modes or 2 full res modes.

- Fixed problem with flash files being added, but not being deleted.

- Add support for new 2520 power supplies (11092).

- Fixed a problem with the configure allowing the user to set 4fps for full res or Rad modes.

- Fixed problem with new motherboards failing to retrieve  Base defect or Aux Defect maps over Ethernet.

- Fixed a problem with recursive filter not being able to be changed will acquiring a fluoro image.

- Added support for VMS Mac (Physical Ethernet) addresses.

- Added Fiber-Optic Power supply to System Configuration window

Appendix B: Known Defects

- Time to change modes is excessive due to slow speed at which data can be written to the Global Control card.

- Base defect maps in radiography are not applied after transmitting the base file.  To get new base map to apply, perform a gain calibration or power cycle command processor.

- Viva doesn’t support normal fluoro modes below 7.5 frames per second.  However, the configuration application and the controller software will support frame rates below 7.5 frames per second.

- After operating in zoom mode, then switching to another mode, the upper and bottom portions of the image may have a slight offset shift then the middle of the image.  This difference in offset will dissipate over the first few seconds.

| RECORD OF REVISIONS |  |  |  |
| --- | --- | --- | --- |
| REV# | ORIGINATOR | REV DATE | CCN #/ COMMENTS |
| A | Keith Gray | 1/22/99 | First released Version |
| B | Keith Gray | 2/3/99 | See section 3 of this document for changes. |
| C | Keith Gray | 2/12/99 | See section 3 of this document for changes. |
| D | Keith Gray | 4/1/99 | See section 3 of this document for changes. |
| E | Keith Gray | 6/1/99 | See section 3 of this document for changes. |
| F | Keith Gray | 6/16/99 | See section 3 of this document for changes. |
| G | Keith Gray | 7/27/99 | See section 3 of this document for changes. |
| H | Keith Gray | 12/20/99 | See section 3 of this document for changes. |
| J | Keith Gray | 4/7/2000 | See section 3 of this document for changes. |
| J3 | Keith Gray | 12/10/2000 | See section 3 of this document for changes. (CCN: 1008559). Input into SAP |
| J4 | Keith Gray | 4/26/2002 | See section 3 of this document for changes. (CCN: 1012325) |
| J5 | Keith Gray | 05/22/2003 | See section 3 of this document for changes. (CCN: 1015004) Issued Material number for 2520 System Software. Added support for 4.1, 2520 DCDS Receptors. |
| J5 Bld66 | Keith Gray | PENDING RELEASE | See section 3 of this document for changes. (CCN: 1022668) Issued Material number for 2520 System Software. Added latest 4.1 Receptor firmware . |